## Slide 1: 第二週新人週報
工號:1200106
姓名:黃劭宸
到職日:109/03/30

## Slide 2: E500 SOURCE HEAD
DISMANTLING
ASSEMBLE
ISOLATION TEST
引用自SAMUEL - 山姆爾的E500 SOURCE HEAD Maintenance Guideline

## Slide 3
DISMANTLING
OF COMPONENTS

## Slide 4
-SPARE for MAINTENANCE-
REMOVE THE RETAINER PIN [4 PCS.]
TAKE-OUT THE ARC SLIT
INSIDE THE ARC CHAMBER
[BURN-OUT FILAMENT]
FILAMENT
1
3
2
把retainer pin往上推起 再向外拉出

## Slide 5
REMOVE 2 SCREWS CONNECTING THE
FILAMENT FEED-THRU and FEED STRAP
FEED STRAP
FILAMENT
FEED-THRU
REMOVE 2 SCREW OF FILAMENT CLAMP
CAREFULLY TAKE-OUT THE FILAMENT CLAMP
ATTACHED ON FILAMENT END
NOW YOU CAN FREELY DE-ATTACHED THE
ARC CHAMBER ASSEMBLY
4
6
5
5
用十字起把連接FILAMENT FEED-THRU和FEED STRAP的兩顆螺絲還有固定FILAMENT CLAMP
的兩顆螺絲轉出
用一字起把夾住FILAMENT 的FILAMENT CLAMP的縫隙撐開後 即可取出

## Slide 6
REMOVE THE SIDE PLATES
REMOVE THE CONSUMABLE PARTS ON
REPELLER END and FILAMENT END
COLLECT THE USED CONSUMABLE PARTS
AND PUT IT ON THE ASSIGN SCRAP BINS.
REMOVE ARC CHAMBER [step-aside]
10
9
8
7

## Slide 7
ASSIGN SCRAP BINS
LOOSEN THE GAS LINE SUS
TAKE-OUT THE BASE PLATE and Gas Feed Tube
REMOVE THE TOP and BOTTOM SUPPORT PLATE
BOTTOM
SUPPORT
TOP
SUPPORT
14
13
12
11

## Slide 8
REMOVE THE FILAMENT CLAMP INSULATOR
REMOVE THE O-RING
CLEAN THE SOURCE BODY USING H202, SCRUB IT WITH DISH SCRUBBER, RINSE WITH D.I. WATER, SPRAY THE SURFACE WITH AIR SPRAY and WIPE WITH IPA… LET IT DRY
OTHER COMPONENTS CLEAN IT ON
SAND BLAST CLEANER
18
17
16
15

## Slide 9
ASSEMBLE
OF COMPONENTS

## Slide 10
E500 SOURCE HEAD BODY
Filament FEED-THRU

## Slide 11
TOP/BOTTOM
SUPPORT PLATE
Inner Face
Outer Face
Install

## Slide 12
Install The
BASE PLATE
Gas Hole
Gas Feed Tube
Filament Feed-Thru Rod (+)
*Notice The Proper Orientation Of
Locator Slot While Installing The
TOP and BOTTOM SUPPORT Plate

## Slide 13
Install Here!
MAKE SURE THAT THE GAS FEED TUBE and GAS FEED-THRU is
CLOSELY ALIGN
BEFORE TIGHTEN SUS
FILAMENT
CLAMP
INSULATOR
TIGTHEN
the SUS
GAS FEED TUBE
GAS FEED-THRU

## Slide 14
BOTTOM END PLATE
BOTTOM SHIELD END PLATE
ASSEMBLE THE FOLLOWING REPELLER END COMPONENTS
CERAMIC INSULATOR
REPELLER
SHORT
LONG

## Slide 15
REPELLER END
Slot For
SIDE PLATES
At Least 3 Lines Of Thread
Protrude (alignment matter)

## Slide 16
ASSEMBLE THE FOLLOWING FILAMENT END PLATE COMPONENTS:
FILAMENT
FILAMENT
SHIELD
FILAMENT
SCREW
FILAMENT INSULATOR
END CAP
PLATE
INSULATOR
NUT

## Slide 17
FILAMENT END PLATE
FILAMENT CLAMP
Install Here!

## Slide 18
ARC CHAMBER SIDE PLATES
Assemble The Arc Chamber Parts
Check If The SIDE PLATES Slot is 1.0 mm DEEP…
If Not or It Cannot Hold In Place Both FILAMENT
and REPELLER SHIELD… Replace New SIDE PLATES

## Slide 19
Here!
Install ARC CHAMBER ASSY.

## Slide 20
ARC CHAMBER
FILAMENT CLAMP
Filament Strap
Double check if the
screws is tighten
Double Check or Make Sure That Both
REPELLER SHIELD and FILAMENT SHIELD
Must Not LOOSE or UNFASTEN on
SIDE PLATES Slot

## Slide 21
Install Filament Clamp
Make sure the FILAMENT is not Protruded on the CLAMP
Secure Tighten
The Screws…

## Slide 22: ISOLATION TEST USING DVM METER

## Slide 23
FILAMENT
to
FILAMENT
FILAMENT Line
to
SOURCE Body
1
2
Must Be OVERLOAD !
Ideal Reading
3
Set to OHM
Check The
REPELLER
INSULATOR
IDEAL READING

## Slide 24
Must Be OVERLOAD !
REPELLER
to
END PLATE
REPELLER
to
SIDE PLATE
Must Be OVERLOAD !
5
4

## Slide 25
ARC SLIT
Install on top and secure with 4 RETAINER PIN
RETAINER PIN
Clean O-RING Slot with IPA

## Slide 26
Clean O-RING…
Double Check for any Scratch or Damage…
Apply a little amount of High Vacuum Grease
then INSTALL…
Install The O-RING
Sign NAME
and DATE
Maintenance Done.